PRECISION SQUARE

# CALIBRATION PROCEDURES # 5

ENVIRONMENT:

The environment must be relatively dust free, clean the surrounding area with Citri-clean, and clean all inspection tools.

A.	VISUAL INSPECTION

1.    Inspect the measuring surfaces for nicks, burrs, and file if necessary.

-Spray toolmakers ink on the measuring surface, let dry – 2 minutes.

-Remove the burrs with a small file.

-Inspect the surface using a straight edge/light penetration

-Remove any high spots.

2.	Inspect the casting body for cracks, nicks.

3.     Inspect the blade for wear, nicks, burrs and readability.

- File or replace if necessary.

4.     Inspect the Clamping Bolt Assembly for excessive wear.

- Replace components as required.

## B.	CLEANING

- Clean all components.

- Use WD-40 for cleaning the components.

- Remove excess WD-40 with a clean, dry, lint free cloth.

- Assemble components.

## C.	INSTRUMENT PERFORMANCE

- Inspect the function of the Clamp Assembly, ensure blade clamps tightly.

- Ensure blade slides smoothly when unclamped.

## CALIBRATION

- Clean Angular Gauge Block – Use alcohol based cleaner and a lint free cloth.

- Clean 90( Precision Square – Use Alcohol based cleaner and a lint free cloth.

- Ensure measuring surfaces on calibration and measuring instruments are clean.

- Measure 90( outside surface – record findings in Inspection Report.

- Measure 90( inside surface – record findings in Inspection Report.

- Measure 45(surface – record findings on Inspection Report.