MITUTOYO SERIES 180 PROTRACTOR CALIBRATION PROCEDURE #4

ENVIRONMENT:

The environment must be relatively dust free.  Clean the surrounding inspection area with Citri-Clean.  Clean all the measuring instruments and gage blocks with an alcohol based cleaner.

## A.	VISUAL INSPECTION

- Inspect the measuring surfaces for nicks and burrs.  Resurface if necessary.

- Spray toolmakers ink on the measuring surfaces, let dry – 2 minutes.

- Remove any burrs with a small file.

- Inspect the surface using a straight edge/light penetration.

- Remove high spots until measuring surface is sufficiently flat.

- Inspect the casting body for cracks and nicks.

- Inspect the blade and protractor plate for wear, nicks and readability.

- Deburr or replace if necessary.

- Inspect the clamping bolt assembly for wear.  Do this for the swivel clamp and blade clamp.

- If wear is excessive, replace components as required.

## B.	CLEANING

- Clean all components.

- Disassemble protractor head.

- Once disassembled use WD-40 to clean the components.

- Remove excess WD-40 and oily residue with a clean, dry, lint free cloth.

- Apply molycoat lubricant to bolt threads and main casting swivel.

- Assemble components.

## C.	INSTRUMENT PERFORMANCE

- Inspect the function of the blade/swivel clamp.  The blade and swivel should clamp tight with moderate pressure being applied to clamp bolts.

- Ensure blade/swivel slide and rotate smoothly when unclamped.

## D.	CALIBRATION

- Ensure all contact surfaces on protractor and gauge instruments are clean.

- Complete upper portion on Measuring Instrument Inspection Report.

- Inspect accuracy of protractor as outlined in lower portion of Inspection Report and document findings.

- If applicable, calculate combined repair costs.

- Clean all gauge instruments and return them to their proper place.

- Clean the surrounding measuring instrument area.